## Slide 1: UAO SpO2 Simulator GUI Training
Kevin Machado-Gamboa
London, October 13, 2020

## Slide 2: Project Origin
Universidad Autonoma de Occidente
The Biomedical Engineering Research Group, GBIO
Cali, Colombia
Notes: Sistema fotovoltaico que genera un poco mas del 15% de la energía que consume

## Slide 3: Photoplethysmography Signal Generation
Reference: https://www.howequipmentworks.com/pulse_oximeter/
Red light emition diode
Infrared light emition diode
Red-Infrared light detector
Pulsating signal
Signal to Simulate

## Slide 4: MARCO TEÓRICO
Saturación de oxígeno en la hemoglobina. [22]
La hemoglobina tiene la particularidad de cambiar de color cuando se encuentra en presencia del oxígeno
El cambio de color es el factor clave aprovechado en los oxímetros
Ley de Beer-Lamber
α = coeficiente de extinción
c = concentración de la muestra
d = al espesor de la muestra
Notes: El concepto hace referencia al número de lugares o uniones que el oxígeno ocupa en la molécula de hemoglobina. Para entender el principio de funcionamiento de la oximetría hay que conocer las propiedades ópticas de la hemoglobina

## Slide 5: MARCO TEÓRICO
Las longitudes de onda entre 650 nm y 950 nm no causan daño a las células del cuerpo
La hemoglobina desoxigenada absorbe más luz roja que luz infrarroja en la franja de longitudes de onda cercanas a 660 nm.
la hemoglobina oxigenada absorbe más luz infrarroja que la luz roja en la franja de longitudes de onda alrededor de los 940 nm
Espectros de absorción para las especies O2Hb, HHb, COHb y MetHb. [15]

## Slide 6: MARCO TEÓRICO
Ondas fotopletismográficas (PPG) causadas por la pulsación arterial. [24]
Curvas de calibración teórica
Relación entre mínimos y máximos para las señales PPG en los componentes rojo e infrarrojo

## Slide 7: GUI Python Structure
Reference: mainWindowPPG.py
Libraries Needed
GUI Development
App Execution
Main Class

## Slide 8: mainWindow Class
Reference: mainWindowPPG.py
mainWindow contains 13 functions:
    1 Initial
    4 for SpO2
    8 for window configuration

## Slide 9: Simulating Photoplethysmography Using Sinusoids
Reference: mainWindowPPG.py

## Slide 10: Spo2sl_change

## Slide 11: _initial

## Slide 12: _update_plot

## Slide 13: _update_plot Functioning
https://doc.qt.io/qtforpython/PySide2/QtCore/QTimer.html

## Slide 14: Buttons Functioning
Connect the button with an associated function

## Slide 15: Reset_Buffers Function
Container variables
All GUI configurations need to be run inside the __init__ function

## Slide 16: Software
Qt Designer
Reference: https://github.com/mherrmann/fbs-tutorial

## Slide 17: Graphical User Interface
References:
https://ieeexplore.ieee.org/document/8564698
https://github.com/LatinAmericanProgramer/UAO_SpO2_Sim

## Slide 18: Contact details
Linkedin - https://www.linkedin.com/in/kevin-machado-gamboa-b958b6167/
Github - https://github.com/kevinmgamboa/
Email: ing.kevin@hotmail.com
Thank you very much